Aurora Plate Loader Guide

Cytek Loader Tutorial: https://www.youtube.com/watch?v=ouwc_Hiqkhg

Start Up

- Make sure that tube is loaded onto Aurora for initial calibration of sample probe.

- Remove tube from holder once the initial calibration is completed and the Cytometer is showing ass “Connected.”

- Pull lever on loader to the front position.

- Turn on loader via the button in the back.

Template Set Up

- Go to “Preferences” tab

  - Under “Sample Delivery Mode” select “Plate”

  - Check the box for “Bubble/Clog Detection”

- Return to “Acquisition” tab.

- Click “Add Plate”

Select appropriate type of plate

- Label plate accordingly

- Click on “Edit” icon

- Select “Groups” tab

  - Select wells that need to be acquired

  - Can add multiple groups to a plate or make entire plate a single group

- Select “Acquisition” tab

  - Set “Worksheet” to appropriate worksheet you would like to use

  - Set “Events to Record” to an appropriate value

  - Set “Storage Gate” to appropriate gate

  - Set “Stopping Volume” to an appropriate vale

We used 110 µL on Friday.

- Select “Loader Settings” tab

  - Select the appropriate direction for sampling by selecting the arrow icons

  - Set “Mix Time” to “4”

  - All other settings should be fine.

Sample Probe Depth Check

- Create a new plate with a sample only in well D1

- Use an empty plate and fill D1 with water or PBS

- Run the sample and look at the probe depth while running

  - If sample probe appears to touch near the bottom of the plate, you are ready to run samples

  - If sample probe is too high:

    - Go to “QC& Setup” side tab

    - Click on “Calibrate SIT”

    - Change the value to a smaller number

    - Run plate again to evaluate new SIT depth setting

Run Plate

- Select a well to record data for

- Select an appropriate acquisition speed

- Click “Record” button and select “Record Well”

- If sample settings seem appropriate select “Record Plate” on “Record” button

Shut Down

- Move the lever to adjust the loader to the off position (Move lever to the back)

- Turn off loader

- Under “ Preferences” Tab switch back to “Tube” under “Sample Delivery Method”

- Wash with Contrad

- Wash with Water

- Perform a “Fluidics Shutdown” and turn off Aurora

Troubleshooting/Notes

- If plate holder is not in position to load a plate onto it click “Eject” button to return it to the proper position.