Cell Sorting

Sample Preparation:

The key for a high speed, high purity sort is to avoid cell aggregation.  Aggregation can clog the nozzle and disturb droplet formation.

Sample buffer: PBS + 1% BSA + 1 mM EDTA.  Protein and Ca/Mg increase aggregation.  Serum has both, so avoid it if cells can survive.

Use 1 mM EDTA, starting from the time harvesting the cells.  This improves the sort efficiency.  On a PBMC sample, without EDTA, efficiency 60%; with EDTA, over 90%.

If you use trypsin to detach the cells, do it as short as possible.  3-5 min range.  Longer time results in aggregation.  Avoid using high concentration of FBS to neutralize.  More protein covering the cells will increase aggregation.  1% FBS is good, 2% max.  Centrifuge gently, with just enough speed to bring cells down.

Too many dead cells and DNA released from them can make the sample sticky.

Sample concentration:

- Adherent cells, 3 million/ml

- Non-adherent smaller cells such as lymphocytes, 10-15 million/ml

- Single cell sorting into 96 well plate, 0.5 million/ml

Sort time: 2 ml / h.  When counting the cells, please include dead cells and debris.  Sorter needs to see them to gate them out.

Keep samples on ice.  Filter samples using Falcon 5 ml tube with strainer cap (352235).  Bring 10 ml sample buffer so I can dilute and rinse.

Collection:

Using 100um nozzle, you will get 3.5 ml of PBS / 1 million cells.

85um nozzle, 2 ml PBS/ 1 million cells.

Can sort into 1.5 ml, 5 ml, 15 ml tubes and 96 well plate.

Collection tubes can be:

BD Falcon 5 ml collection tube (Ref 352058)

BD Falcon 15 ml polystyrene conical tube (Ref 352095)

Buffer in collection tubes, PBS+10% serum.  Coat 2/3 of the tube with it overnight or 1 h in advance.  The purpose is to wet the tube wall and coat it with proteins.  You can use other buffer. Avoid culture medium, because PH is wrong without CO2 supply.